## Slide 1: Confocal laser scanning microscopy
with point scanners (the ‚normal‘ confocal, not the spinning disk)
Ludwig-Maximilians-Universität München Walter-Brendel-Zentrum für experimentelle Medizin (Wbex) and Core Facility Bioimaging at the Biomedical Center

## Slide 2: Note 1
This presentation contains material for a lecture on confocal microscopy. It is not a complete lecture, nor is it suggested that all slides should be shown in a lecture – there is some redundancy. Some slides - like this one - are ‘hidden’ (won’t show in presentation mode).
The ‘notes’ for each slide may contain links to download source, e.g. for the historic images, or other information.
File version 2016-July
Notes: Some slides contain usefull information in this space.

## Slide 3: Note 2
This presentation was published on the German BioImaging Website. http://www.germanbioimaging.org under Workgroup 5 Training of facility users German BioImaging is a network of German microscopists and imaging specialists.
And on the website of the Core Facility Bioimaging at the Biomedical Center (BMC) of the Ludwig-Maximilians-Universität München http://www.bioimaging.bmc.med.lmu.de/learning-teaching

## Slide 4: Copyright notice
Except for the logos or if otherwise noted or obviously from other sources, material in this presentation was created by Steffen Dietzel.
Material from Steffen Dietzel is distributed under the license ‘Creative Commons Attribution Sharealike‘. This license includes permission to use, alter and redistribute this material (not limited to but including to students and colleagues), provided you give appropriate credit to the creator (e.g. by leaving the copyright notice at the right lower edge of the slide) and you distribute any derivative work under the same license. For license details see http://creativecommons.org/licenses/by-sa/4.0
The logos (lower and upper right corners of the slides, title slide) are not covered by this license. They may be redistributed with the slides provided in this presentation, but not in a different context.
If you should need a different license for your purpose, please contact me at dietzel@lmu.de

## Slide 5
the confocal principle
beampath in a confocal laser scanning microscope, scanning, zooming

## Slide 6: confocal, adj.
Having the same focus or foci.
1867 W. Thomson & P. G. Tait Treat. Nat. Philos. §494 Any two confocal homogeneous solid ellipsoids of equal masses produce equal attraction through all space external to both.
(„Con“ means „with“, as in chili con carne)
Oxford English Dictionary:
Notes: The Oxford English Dictionary is online at www.oed.com. A subscription is required.

## Slide 7: Confocal laser scanning microscopy
Clsm has 2 fundamental differences compared to conventional LM:
Out-of-focus light is blocked physically
No image is generated in the microscope itself. It is reconstructed in the computer
Since the eye cannot see anything, all commercial confocals have an additional excitation source for conventional fluorescence microscopy, like a mercury arc Lamp or an LED.

## Slide 8
A point light source (1) is imaged into the specimen (2) (by the objective) in one spot, the focal point.
Because of diffraction, this image in the focal plane (2D) is an Airy disk (and not actually a point).
The focal ‘point’ is then imaged (by the objective) into the intermediate image plane (3), where it is again focused into one spot = Airy disk.
(1)
(2)
(3)

## Slide 9
(1)
(2)
(3)
confocal, adj.: Having the same focus or foci.
1867: W. Thomson & P.G. Tait Treat. Nat. Philos. §494: Any two confocal homogeneous solid ellipsoids of equal masses produce equal attraction through all space external to both.
(„Con“ means „with“, as in chili con carne)

## Slide 10
A point light source (1) is imaged into the specimen (2) (by the objective) in one spot, the focal point.
Because of diffraction, this image in the focal plane (2D) is an Airy disk (and not actually a point).
The focal ‘point’ is then imaged (by the objective) into the intermediate image plane (3), where it is again focused into one spot = Airy disk.
(1)
(3)
(2)

## Slide 11
(3)
The pinhole in this plane (3) only lets through this Airy disk, everything else is blocked.

## Slide 12
The pinhole in this plane (3) only lets through this Airy disk, everything else is blocked.
(3)

## Slide 13
Problem: only one spot can be investigated for intensity

## Slide 14
Problem: only one spot can be investigated for intensity
Solution: scanning this spot over the specimen in 3D

## Slide 15
Notes: Example for a layout of a commercial confocal

## Slide 16: Clsm and 3D
Technically, „confocal“ has nothing to do with 3D.
However, all clsm around have the capability of moving the focal plane through the specimen. Thus a series of images can be generated which can be combined to a 3D image stack.

## Slide 17: Gallery of optical, confocal sections (every 4th section shown)
Notes: Fixed HeLa-cell nucleus transgenic for a Histone H2B-GFP fusion protein (displayed in Cyan) and with a scratch replication label highlighting sites of DNA replication (red).

## Slide 19: Neues confocal Bild?
Bilder aus Praktikum
Single sections: Pinhole 1 Airy disk Pinhole fully open
3D reconstruction: Pinhole 1 Airy disk Pinhole fully open

## Slide 20
confocal resolution and optical sectioning; pinhole size

## Slide 21: The image of a point is not a point, but an Airy pattern

## Slide 22: How the Airy patterns and intensity plots in this presentation were created
COSM PSF (http://cirl.memphis.edu/cosmos.php) was used to produce a PSF image (.wu file), see settings in screenshot left. Reassuringly, the first minimum turned out to be at 217 nm away from the Airy pattern maximum, as it should be according to d=0.61*λ/NA.
The .wu image was opened in Fiji (http://fiji.sc/Fiji) using the Import-Raw function with the following parameters: Type 32-real, Offset 1024, Gap 0, Little-endian and the width, height and number of images as defined in COSM PSF.
In Fiji, the image’s type was changed to 16 bit and it was then saved as Text-Image to be reopened in Excel 2013. In the 2001x2001 pixel image, the position of the maximum gray value of 65535 was labeled in Fiji position 1001;1001 and in Excel in Line 1002, Column ALN.
The numbers of line 1002 were then used to draw the “wide-field” diagrams (blue) in GraphPad Prism. For the confocal-with-closed-pinhole diagram (green), the numbers of this line were squared.
For import in Powerpoint, images of Airy patterns in Fiji were first adjusted to the desired grayscale range and then converted to 8-bit,
For images with the ‘added pixel intensities’ (see right), 80x80 pixel (=nm) sized boxes were defined in Fiji (y-value for all: 961. x-values: 961, 1041, 1121, 1201, 938, 1018, 1098, 1178). Two images were added using the “Calculator Plus” (k1=0.5; k2=0). All 12 ‘pixel-’images were combined in a stack to create the montage shown here. The average intensity was determined using the ‘measure’ command. Values were normalized in Excel.
One exception applies: The image with two neighboring Airy patterns was created by Katharina Nekolla with Matlab code kindly provided by Peng Xi. This code is published in his book “Optical Nanoscopy and Novel Microscopy Techniques" published by CRC.
I would like to thank the contributors to the ‘confocal list’ (http://lists.umn.edu/cgi-bin/wa?A0=CONFOCALMICROSCOPY). Without the tips of Arne Seitz, Sergey Tauger, David Baddeley, Peng Xi and others, I would have not been able to generate those images.

## Slide 23: The image of a point is not a point, but an Airy pattern
Sir George Biddell Airy,
27.07.1801–02.01.1892
Astronomer Royal 1835-1881
Notes: Source of photography: http://wellcomeimages.org/indexplus/image/V0025959.html
Via https://commons.wikimedia.org/wiki/File:George_Biddell_Airy2.jpg

## Slide 24: Intensity profile
measured along this line

## Slide 25: Example: Intensity distribution in the Airy pattern of a point with an NA 1.4 objective with λ=500 nm
217 nm
217 nm

## Slide 26: How close can two Airy patterns be together and still be resolved = recognized as two?
Lord Rayleigh
(John William Strutt)
12.11.1842-30.07.1919
Notes: Source of photography: http://www.nobelprize.org/nobel_prizes/physics/laureates/1904/strutt-facts.html

## Slide 27: Example: Intensity distribution in the Airy pattern of a point with an NA 1.4 objective with λ=500 nm
26%
Rayleigh criterion

## Slide 28
“..This rule is convenient on account of its simplicity; and it is sufficiently accurate in view of the necessary uncertainty as to what exactly is meant by resolution.”
Rayleigh (1879) ‘XXXI. Investigations in optics, with special reference to the spectroscope, Phil. Mag. Series 5,8: 49, p 267
Lord Rayleigh won the Nobel Prize in Physics 1904
"for his investigations of the densities of the most important gases and for his discovery of argon in connection with these studies"
Lord Rayleigh
(John William Strutt)
12.11.1842-30.07.1919

## Slide 29: But how do we measure that?
It is very difficult to exactly determine the location of the minimum experimentally
Therefore, for measurements a different criterion is needed:
The Full Width Half Maximum, FWHM

## Slide 30: Example: Intensity distribution in the Airy pattern of a point with an NA 1.4 objective with λ=500 nm
182 nm
182 nm
FWHM

## Slide 31: Example: Intensity distribution in the Airy pattern of a point with an NA 1.4 objective with λ=500 nm
normal fluorescence
confocal, closed pinhole
217 nm

## Slide 32: Example: Intensity distribution in the Airy pattern of a point with an NA 1.4 objective with λ=500 nm
confocal, closed pinhole
217 nm

## Slide 33: Example: Intensity distribution in the Airy pattern of a point with an NA 1.4 objective with λ=500 nm
confocal, closed pinhole
155 nm
26%

## Slide 34: But:
If we close the pinhole, we don’t get a signal on the detector. Therefore, the pinhole is usually set to 0.8 – 1 AU.
AU = Airy unit, diameter of the first ring-minimum of the Airy pattern

## Slide 35: Confocal resolution and intensity vs Pinhole size
Signal intensity
Resolution (FWHM)
Data source: CJR Sheppard et al. (2006) Signal–to-noise ratio in confocal microscopes. In: JB Pawley (ed) Handbook of biological confocal microscopy (3rd ed), pp 442-452. Numerical data extracted with Engauge Digitizer V.4.1 from Fig 22.2 (circular pinhole) and Fig 22.3 (1/[transverse resolution (point)]).

## Slide 36: All this was about imaging in the focal plane – what about z?

## Slide 37: The Airy pattern in z-direction

## Slide 38: The Airy pattern in z-direction

## Slide 39: Example: Intensity distribution in the Airy pattern of a point with an NA 1.4 objective with λ=500 nm
Z-direction

## Slide 40: Example: Intensity distribution in the Airy pattern of a point with an NA 1.4 objective with λ=500 nm
Z-direction

## Slide 41: Example: Intensity distribution in the Airy pattern of a point with an NA 1.4 objective with λ=500 nm
FWHM(z) 472 nm
Z-direction
Notes: In this graph (computed in CosmPSF with 10 nm steps) the half maximum z is actually at -0.25 µm and 0.25 µm, the FWHM in this graph is therefore 0.49.

## Slide 42: Example: Intensity distribution in the Airy pattern of a point with an NA 1.4 objective with λ=500 nm
Z-direction

## Slide 43: Example: Intensity distribution in the Airy pattern of a point with an NA 1.4 objective with λ=500 nm
Z-direction
theoretical
confocal

## Slide 44: Example: Intensity distribution in the Airy pattern of a point with an NA 1.4 objective with λ=500 nm
Z-direction
theoretical
confocal

## Slide 45
confocal detection of fluorescence and reflection; laser transmission detection
equipment

## Slide 46: Confocal Microscopy
To produce confocal images, in principle no laser or fluorescence is required:
Transmission confocal microscopy is possible (e.g. Minsky, 1957) but not commercially available.
The point is, that only one spot of the preparation is scanned at any one time and only light from that point is detected.

## Slide 47
Detectors: PMTs and hybrid detectors

## Slide 48: Photomultiplier Tube (PMT)
Image created by Jkrieger and put into the public domain.
English Translation by Steffen Dietzel. Image source:
https://commons.wikimedia.org/wiki/File:Photomultiplier_schema_de.png

## Slide 49: PMT
Typically, a PMT has 10 Dynodes. If 5 electrons are sputtered for each hitting electron, amplification is 510, i.e. about 10 million electrons for each photoelectron from the photocathode
As long as the PMT is not saturated, the total number of secondary electrons is generally proportional to the number of hitting photons, but:
Amplification is a statistical process: one electron may sputter 5 electrons, the next one maybe 4 or 6.
This statistical noise (Poisson noise) gets less important when image contrast is increased by longer exposure, e.g. image averaging.

## Slide 50: PMT
The analog signal is then digitized, usually with 8 bit (28 = 256 values, 0-255).
Higher voltage (= gain) leads to a stronger signal and to more electronic noise.
the „offset“ defines the zero-point (how many electrons are „zero“).

## Slide 51: Photomultiplier Tube
If an electron is amplified 3 times on each dynode, and 10 dynodes are present, a total amplification of 310 = 59049 is achieved
However, amplification is a statistical process!
In this example, producing just one electron less or more on the first dynode may cause a signal deviation of 1/3.
Thus, the final numbers of electrons for a given number of photoelectrons are smeared over a large range

## Slide 52: How many photons do we really collect in one pixel?
Let’s have a look at the grayvalues in an image…

## Slide 53: SiR-actin staining, recorded with PMT, no averaging Clipping size 629x548 = 344692 pixels
Formaldehyde fixed HUVEC cells

## Slide 54: SiR-actin staining, recorded with PMT, no averaging Clipping size 629x548 = 344692 pixels
Formaldehyde fixed HUVEC cells
Notes: The histogram shows the number of pixels having a given gray value (y-axis) for each of the 256 gray values (x-axis) of the image

## Slide 55: The same optical section recorded with Hybrid detector (and less laser power)
Formaldehyde fixed HUVEC cells

## Slide 56: The same optical section recorded with Hybrid detector (and less laser power)
Formaldehyde fixed HUVEC cells

## Slide 57: The same optical section recorded with Hybrid detector (and less laser power)
Formaldehyde fixed HUVEC cells

## Slide 58: How many photons do we collect per pixel?
Actual grayvalues
0 11 23 35 47 59 71 83 95 107 119 131 143 155 167 179 191 203 215 227 239 255
Number of Photons
0 1 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 21 >21
Histogram for a confocal image, recorded with PMT and hybrid detector
Number of pixels with this gray value
grayvalue

## Slide 59: Hybrid detector
Notes: More info on hybrid detectors:
http://www.leica-microsystems.com/science-lab/sensors-for-true-confocal-scanning/

## Slide 60: Hybrid detector
A hybrid between a PMT (Vacuum acceleration) and an avalanche photodiode (semi-conductor amplification)
The amplification happens in only two steps and the vacuum acceleration is very high (8 kV). Thus it produces less statistical noise than a PMT:
The number of photons is directly related to the gray level!
A smaller photocathode results in less dark noise.
The vacuum voltage is fixed. No adaption to very bright signals is possible.

## Slide 61: Direct comparison: Single scan of a fluorescent slide PMT HybridDetector
Notes: 64x64 clipping (of a 256x256 image) was scaled up 4x without interpolation in Fiji, to show pixelization. PMT was set to 1000V (of 1250 max), the HyD to “smartGain 100”. Other settings were identical. Recorded on a Leica SP5.

## Slide 62: Histogram of a single 256x256 scan of a fluorescent slide
Chart: Histogram HyDs vs PMTs
PMT HyD

## Slide 63: Histogram of a single 256x256 scan of a fluorescent slide
Chart: Histogram HyDs vs PMTs
Normalization

## Slide 64: Histogram of a single 256x256 scan of a fluorescent slide
Chart: Histogram HyDs vs PMTs
Notes: Normalizaton of PMT-x-values with factor x1.8

## Slide 65: Average Projection of 100 Scans (frame average) PMT HyD
Notes: 64x64 clipping was scaled up 4x without interpolation, to show pixelization. PMT was set to 1000V (of 1250 max), the HyD to “smartGain 100”. Other settings were identical. Recorded on a Leica SP5.

## Slide 66: Histogram of a 100x average 256x256 scan of a fluorescent slide
Chart: 

## Slide 67: Histogram of a 100x average 256x256 scan of a fluorescent slide
Chart: Histogram HyDs vs PMTs
Normalized:
Notes: Normalizaton of PMT-x-values with factor x1.71

## Slide 68
Left: PMT with 7% laser power (633 nm) Right: Hybrid detector with 1%.
Lookup table ‘fire’ from Fiji
Note: this does not translate into “HyD 7x better than PMT”!

## Slide 69: How many photons do we collect per pixel?
Histogram for a confocal image, recorded with PMT and hybrid detector
Number of pixels with this gray value
grayvalue
PMT
Hybrid detector
Notes: Cells were stained with SiR-Actin to visualize the actin sceleton. On a Leica SP5, a 2D image section was recorded twice, first with a PMT (alkali photcathode) and then with a hybrid photodetecor. From these two images, the above shows the distribution of gray values. The actual images can be found in the file SD_Resolution_in_Fluorescence_Microscopy.pptx
The distribution shows that the number of generated photoelectrons (the electron generated by photon impact) per pixel can be directly seen with hybrid detectos but not with PMTs. See also next slide.

## Slide 70: PMTs and hybrid detectors
The spectral response, quantum efficiency, and dark current of a detector depend on the material and size of the photocathode.
The best photocathodes capable of responding to visible light used to have less than 30% quantum efficiency: 70 percent of the photons impacting on the photocathode do not produce a photoelectron and are therefore not detected.
New GaAsP photocathodes have a QE of up to 45%. Used in some PMTs (Zeiss, Nikon) and always in hybrid detectors (Leica)

## Slide 71: Hybrid vs PMT
Hybrid detectors have less dark noise due to smaller photocathode
Less statistical noise due to higher voltage and fewer amplification steps
Voltage cannot be adapted to varying signal intensity
PMTs can be easily adapted to very bright signals by lowering the voltage (Gain).
Photocathode may be GaAsP or Alkali

## Slide 72: PMTs and hybrid detectors...
...are colorblind! All you get is a gray level image.
The „color“ of the photons arriving at the PMT is determined solely by the filters/optics in front of the PMT.

## Slide 73: Saturation of fluorochromes
Fluorochromes already in an excited state cannot be excited again -> with high light intensities, their response is not linear any more
Laser excitation (>1mW) can cause this situation, since all excitation is concentrated to one point (higher peak intensity than with widefield excitation!)
Saturation is a property of the individual fluor molecule: It happens in bright areas as well as in dark ones.
Solution: lower laser power and average

## Slide 74: Special Effects in clsm
Accusto-Optical Filters
Spectral detection
Spectral sensitivity of detectors

## Slide 75
Sources of noise

## Slide 76: Practical tasks:
Recording of a PSF (beads, Q-dots) at the coverslip and at the slide, compare FWHM in x,y and z. Compare various embedding media.
Record and compare PSF in the center and the edge of field of view.
Measure chromatic aberration with 500 nm polychromatic beads.
Biological specimens:
Can more details be distinguished when a pixel size of 50 nm is used instead of 100 nm?
Compare doubled laser intensity vs. double averaging
Compare confocal image AU 1 vs., AU 0.5 vs. opened pinhole

## Slide 77: How to record a good image

## Slide 78: Before you begin: Check your microscope! (or make sure that somebody else does)
FWHM of the PSF as it should be? (measure with the embedding medium that you plan to use!)
Measure chromatic aberration, especially if you want to do colocalization studies (measure with the embedding medium that you plan to use!). For high NA objectives, expect to find up to 500 nm aberration in z-directon.

## Slide 79: Rule # 1: Start with a good preparation!
“Garbage in, garbage out” (anonymous)
Use a 170 µm coverslip. Pay attention to Ri mismatch.
Fixed cells: Don’t destroy the structure of your cells if you want to look at structural organization
Hepperger C, Otten S, von Hase J, Dietzel S. (2007) Preservation of large-scale chromatin structure in FISH experiments. Chromosoma 116(2):117-33. DOI: 10.1007/s00412-006-0084-2
Open Access, © The Authors
Notes: Figure Legend: Dual color FISH on mouse ES cells with a paint probe to chromosome 7 (green), a BAC (red) and DNA counterstain (blue). All micrographs are at the same scale (bar 5 μm). While territories in buffered formaldehyde (BF) fixed nuclei (a, b) appear compact, territories after hypotonic Methanol Acetic Acid fixation (c–e) mostly have a disrupted, torn appearance which is in line with the increased diameter of their nuclei. In BF-fixed nuclei, BAC signals are connected to chromosome territory signals. For nuclei in (a) and (b), a single confocal section (left) and a projection of the stack (right) are shown. (c)–(e) are single confocal sections. (f) Object counting reveals a much higher disintegration of chromosome territories in Hypo-MAA-fixed nuclei (n = 33) than in BF-fixed nuclei (n = 39). While values below a threshold of 80 are influenced by nuclear background, above 200 chromosome territories start to fall below threshold. From 80 to 200, territories in BF-fixed nuclei show separation in about two objects, one for each chromosome territory, while in Hypo-MAA-fixed nuclei, the values are around 6. Also, the variation is much larger, indicated by the larger standard deviation (error bars)

## Slide 80: The Uncertainty Principle of fluorescence based microscopy:
The more detailed you look, the more you destroy (bleaching, phototoxicity).
The highest physical resolution and contrast may require high laser powers and are thus often incompatible with specimen preservation and biological reliability

## Slide 81: The eternal triangle of compromise
High intensity
Good Signal-to-Noise ratio
High spatial resolution
High speed
High temporal resolution
Low bleaching and phototoxicity
Modified after Shotton DM (1995): Robert Feulgen Prize Lecture 1995. Electronic light microscopy: present capabilities and future prospects. Histochem Cell Biol. 104(2):97-137. DOI: 10.1007/BF01451571

## Slide 82: Rule # 2: Find the best possible compromise between image quality and introducing artefacts
Scan speed is often indicated as the number of lines per sec. Examples: 400 Hz, 1200 Hz
Averaging: line (each line is scanned x times and the average is immediately computed, good for moving objects) or frame (full frame image is repeatedly scanned and averaged)
Accumulation (graylevels of subsequent scans are added up; useful for weak signals, best with low-noise detectors)
Voxel size (for best resolution with NA 1.4: ~80 nm in xy, 240 nm in z)

## Slide 83: And then: Check for bleaching!
Record the image stack again with the same settings than before:
Is the second stack (nearly) as bright as the first one?

## Slide 84: Opening the pinhole for brighter images?
You don‘t want just ’bright‘ images, you want a good signal to noise.
What is signal, what is noise? Depends on your sample and question.

## Slide 85: Example 1: Calcium ratio measurement in isolated blood vessels (t-series)
A CFP-YFP Ca-detector (TN-XXL) is distributed in the whole cytoplasm and the signal is weak. Here, opening the pinhole led to an increase in „signal“ because out of focus fluorescence contributes to the „signal“.
For experimental details see: Direnberger, S., M. Mues, V. Micale, C.T. Wotjak, S. Dietzel, M. Schubert, A. Scharr, S. Hassan, C. Wahl-Schott, M. Biel, G. Krishnamoorthy, and O. Griesbeck. 2012. Biocompatibility of a genetically encoded calcium indicator in a transgenic mouse model. Nat Commun. 3:1031.
Notes: Note: For experimental details see indicated publication. The figure itself, however is not part of that publication. © Steffen Dietzel

## Slide 86: Example 2: Small dots (replication labeling) in cell nuclei
No signal-to-noise improvement by opening the pinhole, just more background! Because this „signal“ is limited to the focal plane anyway.
Pinhole open
Pinhole 1 Airy disk

## Slide 87: In most cases, opening the pinhole decreases the signal-to-noise ratio: Only the noise gets brighter.
In most cases, pinhole size of 1 Airy unit or a little less (0.8) is best.

## Slide 88: Rule # 3: Make sure the dynamic range covers all structures of interest in your sample
Do not overexpose!
With 8-bit images, set PMT offset and gain such that your image uses the 256 available graylevels as good as possible.
When recording, use a look-up table (LUT) that will show you when you over- or underexpose in warning colors.

## Slide 89: Bleed through = cross talk
Make sure you don‘t get bleed through!
FITC goat anti–mouse IgG antibody/pH 8.0
Spectrum from http://probes.invitrogen.com/servlets/spectra?fileid=143iggp8
FITC
Cy3
Excitation-
spectrum
Emission-
spectrum
Notes: This slide shows why FITC (and related fluors) must not be excited when the orange channel is recorded.

## Slide 90: Transmission image
If your confocal has a transmission laser detector, you can get a transmission image „for free“, i.e. without additional damage to the sample.
Consider if such an image would help to understand your data.
If yes, adjust your condenser to Köhler Illumination! This can be done by using white light halogen (=normal) transmission light.
Set your laser power to achieve a good fluorescent image first
Then, be prepared to play around with the gain and offset a while before you find the optimal setting for the transmission setting. Try closing and opening the condenser aperture diaphragm for better contrast.

## Slide 91: Reflection confocal image
A reflection confocal image may help to identify some structures, such as pseudopodia of cells.
Set up a normal PMT such that it collects the excitation light of the laser. In theory, this light should not arrive at the detector, due to the dichroic beam splitter. In practice, the small percentage of reflected light passing the dichroic is sufficient to create an image.
Be careful! Too intense light may damage the PMT! Start with low laser power and/or low PMT-voltage and work your way up.
When you have determined the limits of safe operation, set your laser power to achieve a good fluorescent image, before calibrating voltage of the reflection image.
In contrast to the transmission image, the reflected image is a true confocal image, since the light passes the pinhole. Depending on sample type an Ri-differences in the sample, this may or may not lead to useful results.

## Slide 92: Confocal History

## Slide 93: References for ‚History‘
References for the history section can be found in the English (https://en.wikipedia.org/w/index.php?title=Confocal_microscopy&oldid=632060817#History ) or, in more detail, in the German (https://de.wikipedia.org/w/index.php?title=Konfokalmikroskop&oldid=131310175#Geschichte ) Wikipedia article on the topic.

## Slide 94: 1940: Hans Goldmann
The ophthalmologist from Bern, Switzerland, develops a slit lamp system to document eye examinations
This system is considered by some later authors as the first confocal optical system

## Slide 95: 1943: Zyun Koana
Publishes a confocal system in Japanese.
No translation is available, so it is not clear what this system actually is.
A figure in this publication, however, clearly shows a confocal transmission beam path.

## Slide 96: 1951: Hiroto Naora
Describes a confocal microscope for spectrophotometry in the journal Science

## Slide 97: 1955, 1957: Marvin Minsky
The first confocal Scanning-Microscope
Figure from the patent application
Notes: In general, the contents of United States patents published before March 1, 1989 are in the public domain.
http://www.uspto.gov/news/media/ccpubguide.jsp

## Slide 98: Minsky‘s design with epi-illumination and mirror
This is not a reflection confocal design, but transmission with subsequent mirror
Figure from the patent application
Notes: In general, the contents of United States patents published before March 1, 1989 are in the public domain.
http://www.uspto.gov/news/media/ccpubguide.jsp

## Slide 99: 1967: Mojmír Petráň Tandem-Scanning-Microscope
The first confocal with a Nipkow-disk.
The first confocal which was commercially available. And for a long time the only one.
Figure from the patent application
(modified: red line added at place of disk)
Notes: In general, the contents of United States patents published before March 1, 1989 are in the public domain.
http://www.uspto.gov/news/media/ccpubguide.jsp

## Slide 100: 1969: The first confocal Laser-Scanning-Microscope
Egger und Davidovits, Yale University, 1969 and 1971, point scanner.
Epi-Illumination-reflection microscopy for the observation of nerve tissue.
Speculation about fluorescent dyes for in vivo investigations.
They cite Minsky‘s patent, thank Steve Baer for suggesting to use a laser with ‚Minsky‘s microscope‘ and thank Galambos, Hadravsky and Petráň for discussions leading to the development of their microscope.
Motivation for the development: in the Tandem-Scanning-Microscope only a fraction of 1/ 10-7 of the illumination participates in generating the image in the eye piece. Thus, image quality is not sufficient for most biological investigations.

## Slide 101: 1969: The first confocal Laser-Scanning-Microscope
Helium-Neon-Laser, 633 nm, 5 mW, is reflected by a semi-transparent mirror towards the objective.
The objective is a simple lens with a focal length of 8,5 mm. As opposed to all earlier and most later systems, the sample was scanned by movement of this lens (objective scanning), leading to a movement of the focal point.
Reflected light came back to the semitransparent mirror, the transmitted part was focused by another lens on the detection pinhole behind which a photomultiplier tube was placed.
The signal was visualized by a CRT of an oscilloscope, the cathode ray was moved simultaneously with the objective. A special device allowed to make Polaroid photos, three of which were shown in the 1971 publication.

## Slide 102: 1977 – 1985: point scanners with laser and stage scanning
In 1977 Colin J. R. Sheppard and A. Choudhury, Oxford, UK, publish a theoretical analysis of confocal and laser-scanning microscopes. It is probably the first publication using the term „confocal microscope“.
In 1978 the brothers Christoph Cremer and Thomas Cremer in Heidelberg publish a design for a confocal laser-scanning-microscope for fluorescent excitation with electronic autofocus.
In 1978 and 1980, the Oxford-group around Sheppard and Tony Wilson describes a confocal with epi-laser-illumination, stage scanning and PMTs as detectors. The stage could move along the optical axis, allowing optical serial sections.

## Slide 103: 1977 – 1985: point scanner with laser and stage scanning
Fred Brakenhoff and coworkers showed in 1979 that the theoretical advantages of optical sectioning and resolution improvement are indeed achievable. In 1985 the group was the first to publish convincing images with regard to cell biological questions, followed shortly afterwards by other groups.
In 1983 I. J. Cox und Sheppard from the Oxford group published the first work connecting a confocal with a computer.
The first commercial laser scanning microscope, the stage-scanner SOM-25 was offered by Oxford Optoelectronics (after several take-overs acquired by BioRad) starting in 1982. It was based on the design of the Oxford group.

## Slide 104: Starting 1985: Laser point scanners with beam scanning
In the Mid-1980ies, W. B. Amos, J. G. White and colleagues in Cambridge built the first confocal beam scanning microscope. The sample was not moving, but the illumination spot was, allowing faster image acquisition: four images per second with 512 lines each.
Hugely magnified intermediate images, due to a 1-2 m long beam path, allowed to use a conventional iris diaphragm as a ‘pinhole’, with diameters ~ 1 mm.
First micrographs were taken with long-term exposure on film before a digital camera was added. A further improvement allowed zooming into the preparation for the first time.
Zeiss, Leitz and Cambridge Instruments had no interest in a commercial production. The Medical Research Council (MRC) finally sponsored development of a prototype. The design was acquired by Bio-Rad, amended with computer control and commercialized as ‘MRC 500’.
The successor MRC 600 was later the basis for the development of the first two-photon-fluorescent microscope developed 1990 at Cornell University.

## Slide 105: Starting 1985: Laser point scanners with beam scanning
Developments at the University of Stockholm around the same time led to a commercial clsm distributed by the Swedish company Sarastro. The venture was acquired in 1990 by Molecular Dynamics, but the clsm was eventually discontinued.
In Germany, Heidelberg Instruments, founded in 1984, developed a clsm which was initially meant for industrial applications, rather than Biology. This instrument was taken over in 1990 by Leica Lasertechnik.
Zeiss already hat a non-confocal flying-spot laser scanning microscope on the market which was upgraded to a confocal.
A report from 1990, mentioning “some” manufacturers of confocals lists: Sarastro, Technical Instrument, Meridian Instruments, Bio-Rad, Leica, Tracor-Northern und Zeiss.